INSTRUCTIONS FOR USE, CLEANING, AND STORAGE

FOR MANOMETRIC CATHETERS

INSTRUCTIONS FOR USE

- 􀂾Post-procedure manometric catheters are contaminated with body fluid.

- 􀂾Only use catheters that have been cleaned and high-level disinfected or sterilized according to instructions below.

- 􀂾Proper aseptic technique and universal barrier precautions (UBP) must apply.

- 􀂾Inspect catheter for any evidence of mechanical damage or imperfections. Do not use if damaged.

- 􀂾Channels that will not be used for measurements should be filled with water and catheter luer connectors sealed with luer plugs or stopcock.

- 􀂾This catheter must only be used under medical supervision by personnel who have been adequately trained in safe gastro-intestinal intubation, and in the performance of perfused manometric measurements.

CHEMICAL CLEANING / DISINFECTING INSTRUCTIONS

1. Cleaning

- 􀂾Wash catheter in warm water with enzymatic detergent solution.

- 􀂾Flush detergent through all channels including unused channels. Remove all foreign matter. Wipe and clean external surface.

- 􀂾Rinse catheter and flush all channels with clean fresh water to remove detergent. Repeat as needed. Blow-dry all channels with air and wipe excessive water from external surface.

2. Chemical Disinfection (high-level disinfection)

- 􀂾Soak catheter in disinfectant solution [the types commonly used for endoscopes; e.g. Cidex-OPA (Johnson & Johnson), Sporox (Sultan Chemists Inc.), Korsolex Extra (Bode Chemie GmbH & Co.)] according to solution manufacturer’s instructions.

- 􀂾Flush disinfectant through all channels including unused channels. Ensure the entire catheter is submerged in disinfectant and all channels are filled with disinfectant.

Caution: The catheter contains Polyvinyl Chloride (PVC), Polycarbonate, Stainless Steel (If a balloon is attached it also contains Latex or Polyethylene, and Silk sutures). Check with disinfectant manufacturer to ensure their disinfectant can be used safely on the materials listed above.

- 3. Final Rinsing and Blow Drying

  - 􀂾Rinse catheter and flush each channel with sterile water to remove disinfectant. Repeat as needed.

  - 􀂾Ensure all disinfectant is completely removed.

  - 􀂾Blow-dry all channels with air and wipe all water from external surface.

ETO GAS STERILIZATION

- 􀂾Sterilization by Ethylene Oxide gas has been validated effective on these manometric catheters.

- 􀂾Enusure pigtail luer ends are open (remove all luer caps) prior to sterilization.

- 􀂾Sterilize at a temperature of 55oC for a minimum full-cycle of 180 minutes.

- 􀂾Aerate at a temperature of 51-59 oC for a minimum of 12 hours.

STORAGE

- 􀂾Ensure all channels and external surfaces are completely dry.

- 􀂾Store catheter dry in open air. Do not store in a sealed container.

Caution: Do not clean or soak catheter in alcohol.

Do not steam autoclave PVC catheters.